### SEKS-SSR-EV- /Week 5 starting   29 Jan 07

### SM18 Operation Meeting (Tuesdays) 30 Jan  07

Bench status  on Mondy 29-Jan-2007 at  2:10:10 PM

CONFIG :  D 6 - AS 3 -  IRSSS 3 ;

4  Magnet Tests completed () ;

3 Dipoles tests

1 500-series ( 509 TT821 not working and EF corrector not reach ultimate 22 quenches )

Situation from Cryogenics:  J-P Lamboy

Situation  all SSS  : J Halik :

Only 3 Dipoles/1 ARC-SSS/ 3 S5xx series remaining for the ring

General Points:

- New shift schedule for cryo, schedule of OP/ICS

- Power cut occurred Wednesday night

- Eric, Kevin, GHH, Florence, Jerome  absent 2 Feb to 18 Feb 2007

- No Test Coordinator. Stéphane will coordinate till end of test.

- MM: MB2267, MB2326. SSS374 and SSS529.

## Tests

| Magnets under Test from 21-01-2007 to 27-01-2007 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR2384 |  |  | 10-01-2007 | 109 h. | 31 h. | 26 h. | 18 h. | 96 h. | 22-01-2007 | 280 h. | Stripping |  |  |
| 2 | TBD2 | SSS509 |  |  | 21-12-2006 | 497 h. | 25 h. | 133 h. | 19 h. | 121 h. | 23-01-2007 | 795 h. | Stand By |  |  |
| 3 | TBE1 | MBAR2414 |  |  | 12-01-2007 | 125 h. | 21 h. | 90 h. | 21 h. | 54 h. | 25-01-2007 | 311 h. | Stripping |  |  |
| 4 | TBF2 | MBAR2408 |  |  | 16-01-2007 | 64 h. | 72 h. | 14 h. | 36 h. | 29 h. | 25-01-2007 | 215 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2384 | 11828.5 A. (D1_U) | 12644.4 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | SSS509 | 11937.3 A. (QI_34) | 50.6 A. (QI 12+34) | 11853.4 A. (QI_34) | 12222.8 A. (QI_12) | 12766.9 A. (QE_34) | 12850 A. | - | - | - | - | - | - | - |
| 3 | MBAR2414 | 11176.9 A. (D2_U) | 11642.8 A. (D2_U) | 12214.5 A. (D1_L) | 12445.7 A. (D2_L) | 12613.2 A. (D1_U) | 12721 A. (D1_U) | 12850 A. | - | - | - | - | - | - |
| 4 | MBAR2408 | 10398.6 A. (D1_U) | 12098.7 A. (D2_U) | 12559 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |

TERMINATION COUNTER - 2007

| No | MAGNET | BENCH | ARRIVAL | DEPARTURE | Comments |
| --- | --- | --- | --- | --- | --- |
| 1 | 2414 | E1 | 12-12-2006 | 25-01-2007 |  |
| 2 | 2408 | F2 | 16-01-2007 | 25-01-2007 |  |
| 3 | 2423 | E2 | 19-01-2007 |  |  |
| 4 | 2328 | F1 | 19-01-2007 |  |  |
| 5 | 2425 |  |  |  |  |
| 6 | 2424 |  |  |  |  |
| 7 | 2326 (MM) | A2 |  |  |  |
| 8 | 2267 (MM) | F2 |  |  |  |
| 9 | 2263 | A1 | 22-01-2007 |  |  |
| 10 | 2426 |  |  |  |  |
| 11 | 2394 |  |  |  | NC on capillary “might” come again |
| 12 | 2790 |  |  |  | Spare |
| 13 | 2624 |  |  |  | Spare |
| 14 | 2598 |  |  |  | Spare |
| 15 | 2420 |  |  |  | Spare |
| 16 | 2419 |  |  |  | Spare |
| 17 | 2418 |  |  |  | Spare |
| 18 | 2399 |  |  |  | Spare |
| 19 | 2551 |  |  |  | Spare |
| 20 | 2739 |  |  |  | Spare |

12 of  20 Dipoles for completion of Tunnel                            Updated : February 13, 2007

Note :

1)   The red line shows the completion of magnets for tunnel.

2)   MBBR 2394 is under investigation for HV problem at Cold (Richard) .  It will be replaced by a spare if this problem could’nt be solved.

3)   One of  xxxx will replace 2324 ?

SICK DIPOLES

| No | MAGNET | Problem due to |
| --- | --- | --- |
| 1 | 2394 | HV at Cold |
| 2 | 1334 | CDBP: QH Vs Gnd |

TERMINATION COUNTER - 2007

| 23 | SSS 523 | D1 | 10-01-2007 |  |  |
| --- | --- | --- | --- | --- | --- |
| 24 | SSS 531 |  |  |  | 19-01-2007* |
| 25 | SSS 529 (MM) |  |  |  | 24-01-2007* |
| 26 | SSS 530 |  |  |  | 01-02-2007* |
| 27 | SSS 361 | B1 | 16-01-2007 |  |  |
| 28 | SSS 356 | C2 | 25-01-2007 |  | 23-01-2007* |
| 29 | SSS 374 (MM) |  |  |  | 26-01-2007* |

7 of  7 SSSs for completion of Tunnel                                    Updated : February 13, 2007

* Expected Arrival Date ( Bldg 904 information)

SICK SSS

| No | MAGNET | Problem due to |
| --- | --- | --- |
| 1 |  |  |
| 2 |  |  |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Wednesday_24_M | TBB1 | SSS361: While launching the MM2 power supply communication error has occurred. | MTM |  | Guy informed, he will check it & also he will re- launch the test |
| Wednesday_24_A | TBF1 | MB2328: MB detrained during second training by 324 Amp. | MTM |  | As per the de-training rules since value is still >12000 A so we can follow the normal training rule chart. Hence PT8, PT9 and then warm up quench. |
| Thursday_25_M | TBB1 | Could not launch MM. MMP Error message "You need a power supply!" The PLCSE on all clusters not working. | MTM |  | Guy informed about MM problems. He found that the water PLC went to standby mode. Solved. |
| Thursday_25_M | TBF1 | MB2328: OP team tried to reset TEMA on F1. Error message "...... UDF socket" | AB/CO | TEMA | Giorgio informed about TEMA problem.Solved |
| Thursday_25_M | All | Problem with e-logbook and e-traveller (Jurek could not sign it). |  |  | Solved by J.M Elyn (e-logbook) and Giorgio’ |
| Friday_26_A | TBA2 | MB2326: Powering of CD EF corrector in the dipole. Inductance value missing. | MTM |  | Stéphane and Guy contacted. Giorgio will power the correctors Monday. |
| Friday_26_A | TBE1 | MB2424: During LF heater discharge (IAP at warm), VT_anode, VT_cathode errors (-1 V). | MTM |  | Salem found that : the connector were connected wrongly the dessicant bag was present in the IFS box. |
| Friday_26_A | TBF2 | MB2267: Magnet connected by MTM team. No WP04 done. | MTM |  | Stéphane asked to go on and check fater the connection the electrical integrity. |